Watermeter Test Manual
# Overview
The Xerofill watermeter test bench provides measured water flows through a watermeter under test. Sensors are provided for water pressure and temperature. A solenoid valve arrangement allows for the water flows to be started and stopped at selected times
This manual describes the components, test setup and measurement procedures for calibration of a watermeter.
# Components
The test bench includes a water flow system incorporating a solenoid valve controlled by software.
The watermeter needs to have special test firmware loaded to count water flow and transmit the results serially via the test connector on board. An interface card converts the serially transmitted signals to RS232 for passing to an external PC. The card also gives access to the raw count signal and allows power to be passed to the watermeter. It provides has a connector for uploading firmware.
An Arduino board with a sensor adapter card performs a separate water flow measurement using the watermeter raw count signal. It also collects pressure and temperature measurements, the latter through an adapter board. The results are transmitted serially to an external PC
A program for the external PC is provided to collect, display and store the incoming data from the two serial sources, as well as control the water flow solenoid.
## Interface Board
Two versions of the interface board are provided. One is free standing for lab testing, while the other is mounted on the testbench.
- The interface boards have a 10 pin header for connecting to the 10 pin header on the watermeter board. The interface board header is polarised and has the larger 0.1” spacing widely used in electronic equipment while the watermeter board header is unpolarised and has the smaller 2mm spacing. The cable connector has a red wire which must be oriented as shown above.
- The free standing interface board has a 2 pin polarised connector for connecting to the power, while the mounted interface board uses a 2.1” barrel socket. This powers both the interface board and the watermeter board.
- The power to the boards must be 5V, with an absolute maximum of 6V, to avoid damage to the integrated circuit.
- A three pin polarised connector is provided for connecting the digitised count signal from the watermeter board to the Arduino.
- A four pin polarised connector is provided for the serial signals sent from the AVR processor on the watermeter board. These will have been converted to RS232 levels by the integrated circuit on the interface board. A Serial-USB adaptor cable will normally required as most PCs now do not have serial interfaces.
- A six pin polarised connector is also provided for uploading firmware to the AVR processor on the watermeter board. This is described further below.
Free Standing Interface Card with 10-pin connector.
Note the orientation of the cable, the red wire appearing at the bottom.
Top View of the Mounted Interface Card
Bottom View of the Mounted Interface Card

## Arduino Data Capture
The
## Temperature Sensor Interface Board
## Watermeter Firmware
# Setup Procedure
# User Interface